APPENDIX 1. LA-ICP-MS instrumentation and Operational setting (Element- XR)

| Laboratory & Sample Preparation |  |  |  |  |
| --- | --- | --- | --- | --- |
| Laboratory name | Dept of Earth Science, ETH Zurich |  |  |  |
| Sample type/mineral | columbite |  |  |  |
| Sample preparation | Conventional mineral separation, 1 inch resin mount, 1um polish to finish |  |  |  |
| Imaging | CL, Jeol 5000, 10nA, 15mm working distance |  |  |  |
| Laser ablation system |  |  |  |  |
| Make, Model & type | Resonetics Resolution 155 |  |  |  |
| Ablation cell & volume | Laurin Technics 155, constant geometry, aerosol dispersion volume < 1 cm3 |  |  |  |
| Laser wavelength (nm) | 193 nm |  |  |  |
| Pulse width (ns) | 25 ns |  |  |  |
| Fluence (J.cm-2) | 2.0 J.cm-2 |  |  |  |
| Repetition rate (Hz) | 5Hz |  |  |  |
| Spot size (um) | 43 um |  |  |  |
| Sampling mode / pattern | Single hole drilling, 5 cleaning pulses |  |  |  |
| Carrier gas | 100% He, Ar make-up gas combined inside ablation cell funnel. |  |  |  |
| Ablation duration (secs) | 40 secs |  |  |  |
| Cell carrier gas flow (l/min) | 0.7l/min |  |  |  |
| ICP-MS Instrument |  |  |  |  |
| Make, Model & type |  |  | Thermo Element XR SF-ICP-MS |  |
| Sample introduction | Ablation aerosol only, squid aerosol homogenization device |  |  |  |
| RF power (W) | 1500W |  |  |  |
| Make-up gas flow (l/min) | 0.95l/min Ar |  |  |  |
| Detection system | Single detector triple mode SEM, analog, Faraday |  |  |  |
| Masses measured | 204, 206, 207, 208, 232, 235, 238 and trace elements |  |  |  |
| Integration time per peak (ms) | 12 ms (mass 204), 20 ms (masses 208, 232, 235, 238), 40 ms (masses 206, 207) |  |  |  |
| Total integration time per reading (secs) | 0.202 sec |  |  |  |
| IC Dead time (ns) | 8 ns |  |  |  |
| Data Processing |  |  |  |  |
| Gas blank | 10 second prior to each ablation spot |  |  |  |
| Calibration strategy | X-36 used as primary reference material, NIST 610 for tracing |  |  |  |
| Reference Material info | X-36 columbite, (von Quadt et al. 2019), in-house SRM material, all isotope ratios are determined by ID-TIMS method |  |  |  |
| Data processing package used / Correction for LIEF | Iolite 3.5 with VizualAge |  |  |  |